## Slide 1
Magnetic Micro-Calorimeters
for Baby-IAXO and
hi-res x-ray spectroscopy
Loredana Gastaldo, Andreas Fleischmann and a large team at
Heidelberg University

## Slide 2: micro-calorimeters
Operation at low temperature (T < 0.1 K):
small specific heat
large temperature change
small thermal noise !
Temperature change
Relaxation to bath temperature
Thermal bath (T < 100mK)
Weak thermal link G
Absorber C
Thermometer
X-ray photon or particle
Thermal detectors !
Somewhat numb: Convert good form of energy to heat!

## Slide 3: metallic magnetic calorimeters
paramagnetic sensor: Au:Er500ppm , Ag:Er
main differences to resistive sensors like TES and thermistors
no dissipation in the sensor
no galvanic contact to the sensor
M and C of weakly interacting spins well understood
signal size:
numerical optimization
non-linearity: small & understood from thermodynamics
important for large arrays
t
T
~1 mK
20 mK
~1 ms
M
T

## Slide 4: detector geometries
planar paramagnetic T-sensor
superconducting pickup loop with persistent current to generate bias B-field
250 mm
meander-shaped pick-up coil
stripline (sandwich) design
filling factor ≈ 0.4
filling factor ≈ 0.8

## Slide 5: two-stage dc-SQUID setup
300K
main advantages:
low readout noise
large system bandwidth (several MHz)
low power dissipation (close to the MMC)
linearization by flux locked loop
< 100 mK
input coil should
match pickup coil

## Slide 6
microwave SQUID multiplexer (µMUX)
readout 100s of pixels with only one HEMT amplifier and two coaxes

## Slide 7: micro-fabricated devices: first 1×8 array for soft x-rays (a decade ago)
1×8 x-ray absorbers
  250μm×250 μm gold, 5 μm thick
  98% qu.-eff. @ 6 keV
  electroplated into photoresist mold
  mech/therm contact to sensor by stems to prevent loss of initially hot phonons
Au:166Er300ppm temperature sensors
Superconducting meander shaped pickup coils
  2.5 μm wide Nb lines
  Ic ≈ 100mA
On-chip persistent current switch (AuPd)
250 mm
Au:Er
on
Nb meander

## Slide 8: signal rise in compact x-ray detectors
rise time: 90 ns @ 30 mK, as expected for the electron-spin-relaxation of Er in Au

## Slide 9: signal decay
decay time
adjusted by metallic thermal link (Au)
here: 3 ms @ 30 mK
On-chip thermal bath
magnetic flux
time

## Slide 10: first 1×8 array for soft x-rays
DEFWHM = 1.6 eV @ 6 keV
DEFWHM = 1.5 eV @ 0 keV
55Mn
Defines state-of-the-art together with TES from NIST and NASA

## Slide 11
8×8 arrays of maXs-1/20/30/200
1.2/2/4/8 mm
Dry dilution fridge
(Bluefors, FI)
8 × 8 pixels for photons up to 1/20/30/200 keV with DEFWHM = 0.5/2/5/30 eV
highly symmetric design -> cross-talk

## Slide 12
maXs-30 an 8×8 pixel array for x-rays up to 30 keV
Absorbers:
each 0.5mm × 0.5 mm
30-50 mm thick gold
9.8 eV
Energy resolution ΔEFWHM = 9.8 eV @ 59 keV
Resolving power: 6000 (limited by gain drifts)

## Slide 13: maXs-30: calibration
non-linearity
non-linearity as thermodynamically expected!

## Slide 14: maXs-30: present calibration residuals
Most lines from literature have too large uncertainty.

## Slide 15: maXs-30: present calibration residuals
Sub-eV agreement for carefully selected calibration lines.
Left-over residuals due to ADCs  sparse set of calib.-lines should allow for ppm-level (FS) calibration

## Slide 16: MMCs get more and more mobile
here @ entrance to Frejus tunnel between Italy and France

## Slide 17
Two maXs-100 systems @ e-cooler of CRYRING @ GSI/FAIR: He-like Uranium

## Slide 18
maXs-30 @ -beamline, PSI: muonic Li, B, C

## Slide 19
maXs-30 @ -beamline, PSI: muonic Li, B, C

## Slide 20
maXs-30 @ -beamline, PSI: muonic Li, B, C

## Slide 21
Setup challenges:
Extremely low background of
Good mechanical stability for operating while moving/vibrating
Good magnetic shielding
Cryostat challenges:
Reliable operation @
Tiltable to 30°
Detector challenges:
High efficiency in energy range
High energy resolution
Active area matching IAXO optics focal spot size
ü
ü
ü
MMCs for IAXO: Challenges
ü
ü
ü

## Slide 22
maXs-30 set-up for IAXO
Design suitable for IAXO telescope
High purity materials for background reduction
D. Unger et al., JINST 16 (2021) P06006, arXiv:2010.15348 [physics.ins-det]

## Slide 23
maXs-30 set-up for IAXO
55Fe calibration source
Stopping power @10 keV ~100%
Homogeneous performance over the array
Stable operation over 1 month
ΔEFWHM  7 eV
D. Unger et al., JINST 16 (2021) P06006, arXiv:2010.15348 [physics.ins-det]

## Slide 24
maXs-30 set-up for IAXO
D. Unger et al., JINST 16 (2021) P06006, arXiv:2010.15348 [physics.ins-det]
Background spectrum (one month, overground, no special shielding, no muon veto)
Background (1-10 keV): Still too high (target: )
 cryogenic muon-veto to suppress background from muon secondaries

## Slide 25
From maXs-30 to maXs-IAXO
maXs-30
maXs-IAXO
30 µm
10 µm
active area4 × 4 mm2 → 10 × 10 mm2
larger than focus spot
→ intrinsic background monitoring
thickness:30m → 10m to keep ΔEFWHM < 10 eV
Same design, but scaled up!
Axion production
Metallicity
14.4 keV iron line

## Slide 26: summary & outlook
metallic magnetic calorimeters combine in a unique way
energy calibration: sparse set of calibration lines will allow for ppm-FS-accuracy for line-positions in fullrange
micro-fabrication works
scaling to pixelcounts > 100 in progress
large variety of sizes, stopping powers and resolutions
MMCs start to contribute to atomic, molecular, nuclear and astro-particle physics
time resolution
energy resolution
linearity

## Slide 27: MOCCA: 4k-pixel molecule camera
to study reaction cross sections of molecular ions of cold interstellar plasmas
in the cryogenic storage ring CSR at MPI-K, HD
MOCCA:
44.8mm × 44.8mm
64 × 64 pixels
resolution: DE = 250 eV (FWHM)
operated at T = 20 mK
16 two-stage SQUIDs to read-out 64 rows
16 two-stage SQUIDs to read-out 64 columns
Heidelberg University, MPI-K Heidelberg, Columbia Astro Lab NY
Position and energy resolved detection
of neutral molecular fragments
after neutralizing with an electron in the e-cooler
T < 10K

## Slide 28: MOCCA: 4k-pixel molecule camera / hi-res alpha detector
45 mm
200 eV FWHM

## Slide 29: Detector setup V2
New features:
Modularized detector and SQUID holder → easy repair
Optimized for new detector maXs-IAXO
Additional inner teflon shield to reduce background

## Slide 30: thermodynamical properties of Au:Er, Ag:Er
Au:Er 300 ppm
Inverse Temperature T1 [K 1]
Magnetization M [A/m]
Au:Er 300 ppm
Temperature T [mK]
Specific heat C [104 J mol1K1]
Au:Er 300 ppm
signal size can be predicted with confidence
non-linearity as well !
RKKY-interaction between Er-spins ~ 3 times stronger in silver
electrons
thermodynamic properties of interacting spins are well understood
typical range of operation

## Slide 31: MMC-option for LynX: a 100k pixel X-ray camera for astro-physics
Enhanced array hydra:
Main & Enhanced array
absorbers
Main array sandwich hydra:
Main array meander
hydra:
Main & UHR array absorbers
Detector array
4” wafer detector chip
“Large-scale metallic magnetic calorimeter arrays for the Lynx X-ray Microcalorimeter”, Devasia et al. ’22
                read-out by 4000 MUX-ed T-sensors (MMC)

## Slide 32
maXs-30
maXs-30 mounted on coldfinger of a dry dilution fridge

## Slide 33: maXs-30
T

## Slide 34: maXs-30