## Slide 1: Microstrip detector for online monitoring during synchrotron microbeam radiation therapy: measurements with phantoms and veterinary patients
Jayde Livingstone
Equipe Physique Nucléaire et Applications Médicales, LPSC, Grenoble
Jean-François ADAM, Denis DAUVERGNE, Francesca DI FRANCO, Laurent GALLIN-MARTEL,
Marie-Laure GALLIN-MARTEL, Mostafa GHAFOORYAN-SANGCHOOLI, Mohamed IKHELIF, Sarvenaz KESHMIRI, Jayde LIVINGSTONE, Candice MILEWSKI, Jean-François MOTTE, Jean-François MURAZ, Paolo PELLICIOLI, Nicolas ROSUEL, Raphaël SERDUC
F. di Franco, et al., J. Synchrotron Rad., 2023.

## Slide 2: What is radiotherapy?
Use of radiation (e.g. x-rays, electrons, protons, etc.) to treat a disease, such as cancer
About half of all cancer patients will receive some form of radiotherapy during their treatment
The goal of radiotherapy is of course to kill cancer cells…But not at the expense of healthy cells!
Adapted from Thariat J., et al., Lancet, 22(12), 2021

## Slide 3: Microbeam radiotherapy: spatial fractionation of the dose
Curtis, H. J., Radiat. Res., Suppl. 7., 1967
Dose volume effect
http://mswebs.naist.jp
www.esrf.eu
Up until now has been confined to synchrotron sources due to spectrum (kV) and high dose rates (102 – 103 Gy/s) required
Dedicated beamlines:
ID17, ESRF (until mid-2023)
IMBL, Australian Synchrotron

## Slide 4: Experimental Mrt dosimetry
µm beamlets
high dose gradients
high dose rate
keV spectrum
1D dose profile
Si strip detector (10 µm thick SV)
University of Wollongong, Australia
PTW microdiamond (1 µm thick SV)
J. Livingstone, et al., Med. Phys., 2016.
Optical fibres (10 µm thick scintillator)
J. Archer, et al., Sci. Rep., 2019.
2D dose distribution
Radiochromic film
J. Crosbie, et al., Phys. Med. Biol., 2008.
S. Bartzsch, et al., Med. Phys., 2015.
A. Ocadiz, et al., Phys. Med., 2019.
And others….
3D dose distribution
Gel dosimeters (PRESAGE)
S. Doran, et al., Phys. Med. Biol., 2013.
F. Gagliardi, et al., Med. Phys., 2015.
And others…

## Slide 5: In Vivo Dosimetry: Portal imaging
Prescription
Reconstructed dose
Agreement / Disagreement
Dose

## Slide 6: Diamond-based detectors
Diamond satisfies requirements (rapidity, tissue equivalence, radiation resistance)
Operated as a solid ionisation chamber
  Front and back metal electrodes deposited via laser lithography
Induced current due to motion of charge carriers in vicinity of electrodes
  Shockley-Ramo
I. Zahradnik, PhD Thesis, 2021.
Notes: Polarisation effects non-existant when photons traverse
€2000 monocristallin Element6
2x moins cher pour polycristallin DDK
Almax easylab Belgique – thinning, polishing
150 = 550 coupé en 2
Lithogragraphie laser metalisation (Al) 100 nm epaisseur
If less – risk of ripping it off during bonding
Higher precision compared to mask

## Slide 7: Preliminary tests: Dose rate response
2x monocrystalline CVD diamond detectors 4.5 x 4.5 mm² with Al metallisation
  150 µm thick
  550 µm thick
Irradiated at ID17 (ESRF) to investigate dose rate dependence in the range
1 – 104 Gy/s
~100% charge collection efficiency measured using 241Am 5.5 MeV alpha particles (M. L. Gallin-Martel, et al., Frontiers in Physics, 2021.)

## Slide 8: Preliminary tests: Dose rate response
160 kV x-ray tube, 1 Gy/min
95 keV monochromatic x-rays, 0.5 Gy/s

## Slide 9: LPSC microstrip diamond detector
LPSC 8-microstrip diamond detector prototype. All measurements in mm.
8-strip prototype device
550 µm thick monocrystalline CVD diamond
8 strips: 3 mm high, 178 µm wide, 60 µm between adjacent strips
100 nm thick Al metallisation layers
1 V/µm electric field
Guard-ring surrounding strips
Simultaneous integrated charge readout
Aluminium housing with aluminised mylar beam window

## Slide 10: Detector readout
ASIC featuring a charge to digital converter (QDC)
  Coarse measurement
DAC accounts for residual charge
  Fine measurement
Integration time 1 ms – 100 ms
High dynamic range
  10 nA – 100 µA
Vint
Vgr
ASIC QDC
L. Gallin-Martel, LPSC

## Slide 11: Detector response to single microbeam
A single microbeam (as visualised using Gafchromic ® film, left) was scanned horizontally across the microstrip detector.

## Slide 12: Step phantom: measurements vs Theory
In-beam (Eave = 123.7 keV)
Inter-beam (Eave = 308.0 keV)

## Slide 13: 153-Strip detector

## Slide 14: 153-Strip detector
Juxtaposition of 9 x 17-strip diamond detectors:
150 µm thick monocrystalline diamond substrate

## Slide 15: 153-Strip detector response

## Slide 16: Preclinical and veterinary trials (MRT)
Detector
Preclinical trials (rodents):
Glioblastoma
Epilepsy
Veterinary trials (pet dogs):
Brain tumours

## Slide 17: Conclusions
Dose rate linearity over 1 – 104 Gy/s
Homogeneous strip response with no charge loss
Values of µ measured from transmitted beam in good agreement with theoretical values
Next steps:
Analysis of preclinical and veterinary trial data
Full characterisation of 153-strip detector at IMBL, Australian Synchrotron
Development of an upstream detector and dose reconstruction method