## Slide 1
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC
X-ray fluorescence measurements for Pu accountancy in Spent Nuclear Fuel using hard X-ray reflective optics
  Jaime Ruz1,2, M-A Descalle2, K. Schmitt3, J.M. Sharp3 and K. Ziock3
  1Centro de Astropartículas y Física de Altas Energías, Zaragoza, Spain
  2Lawrence Livermore National Laboratory, California, US
  3Oak Ridge National Laboratory, Oak Ridge, Tennessee, US
X-ray DEtector Technologies for Physics. XDEP 2024

## Slide 2
Outline
Introduction
Experimental Setup
  Facility
  Detector
  Reflective optics
Simulations
Measurements of Spent Nuclear Fuel
  GeGI detector
    MOX
  Point contact coaxial HPGe
    MOX
    Bonus run, WAD-125 (considerably younger fuel)
XDEP 2024. J. Ruz, February 6th 2024
1/30
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 3: Spent nuclear fuel represents a global challenge
We are exploring x-ray fluorescence measurements to determine Pu/U ratios
  It contains Pu which is of concern for safeguards
  It is highly radioactive, making traditional NDA analysis challenging
    Particularly true for weak lines
    Pu content inferred from burnup and model calculations
  Reprocessing Plants
    Receiver-operator differences occur on changing from item accountancy and model-inferred mass to destructive analysis of actual material present
For ‘permanent’ disposal facilities (e.g. Onkalo, Finland)
  Desirable to measure content before closing MC&A books
Introduction
Motivation
XDEP 2024. J. Ruz, February 6th 2024
2/30
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 4: K-shell, X-ray fluorescence, we want Pu Ka1:U Ka1
  The Good:
    X-ray emissions are stimulated by high-energy radiation ejecting a K-shell electron
      > 90% of the time one gets K-shell X-ray emission
      These sit on top of the background continuum
  The Bad
    Pu:U ratio determined by elemental abundances ~ 1%
      High background means very tight collimation and long dwell times
  The Ugly
    Pu Ka1 line (103.7 keV) is just 1.6 keV from 155Eu nuclear line (105.3 keV)
      Requires excellent energy resolution
      Incommensurate with high event rates
Introduction
Motivation
XDEP 2024. J. Ruz, February 6th 2024
3/30
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 5
Introduction
4/30
Our ultimate goal: Directly and non-destructively measuring of Uranium and Plutonium content in spent fuel rods.
  Light water reactor spent fuel contains approximately 1% Plutonium and 3% fission products with the remainder Uranium.
  Fission products produce an intense gamma-ray field that obscures the Plutonium gamma rays
    High count rates
    Low intensity of nuclear lines with respect to other radioisotopes
Isotope | Line Energy [keV] | Activity [γ g-1s-1] |  | Isotope | Line Energy [keV] | Activity [γ g-1s-1]
235U | 143.8 | 8.4 x 103 |  | 239Pu | 413.7 | 3.42 x 104
235U | 185.7 | 4.32 x 104 |  | 240Pu | 160.3 | 3.37 x 104
239Pu | 129.3 | 1.44 x 105 |  | 241Pu | 148.6 | 7.15 x 106
239Pu | 375.0 | 3.61 x 104 |  | 241Pu | 208.0 | 2.04 x 107
Research objective
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 6
Introduction
5/30
The radioactive decay of the spent fuel also induces fluorescence in the Uranium and Plutonium and produce x-rays ranging from about 94 to 120 keV.
    These x-rays might prove to be useful NDA signature of the Plutonium to Uranium ratio in the spent fuel.
    The Uranium and Plutonium fluorescence have only rarely been measured in solid spent fuel. Need of optimized and collimated detector system
Element U | Line Energy [keV] | Relative Intensity |  | Element Pu | Line Energy [keV] | Relative Intensity
Kα1 | 98.44 | 100 |  | Kα1 | 103.76 | 100
Kα2 | 94.67 | 61.9 |  | Kα2 | 99.55 | 62.5
Kβ1 | 111.30 | 22.0 |  | Kβ1 | 117.26 | 22.2
Kβ2 | 114.50 | 12.3 |  | Kβ2 | 120.60 | 12.5
Kβ3 | 110.41 | 11.6 |  | Kβ3 | 116.27 | 11.7
Research objective
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 7: Light-water reactor sample spectrum
  30.4 MWd/KgU, 21 months since discharge
  250 mm slit, point contact HPGe, Pu Ka1 is visible, but. . .
MIRROR @ 3.8% deadtime, 2 hours livetime
DIRECT @ 54% deadtime, 2 hours realtime
U Ka1
U Ka2
Pu Ka1
155Eu
U Kb3
U Kb1
U Kb2
Introduction
Candy Slide
6/30
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 8
Use multilayer mirrors to selectively pass emission lines unique to Uranium and Plutonium (and filter out most of the other radiation)
Alternating layers of high- and low-Z materials act as reflecting interfaces, following Bragg’s law
For proof of concept we have deposit multilayers coatings of WC/SiC using constant-d spacing (N=300 , d=1.51 nm) on flat surface of 150x150 mm2
Substrate representation
The tool
7/30
X-ray mirror
Mirror assembly after collimator in a cell wall
Compton-imaging HP Germanium detector (PHDS GeGI) and shielding
Different fluorescence lines will appear at different locations of the detector
Tilting the optic will displace the image of the fluorescence over the face of the detector
m
θ
θ
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 9
Detailed laboratory-based campaign conducted at the Irradiated Fuels Examination Laboratory (IFEL) at Oak Ridge National Laboratory.
The reflective optics were mounted downstream from the 30-inch-long lead collimator with an opening of 5mm x 69.85mm centered on the fuel rod. The 76-cm-long lead collimator was installed in the hot-cell port. An extra collimator with variable aperture was set between the optics and the lead collimation to minimize flux reaching the experimental setup. We also made use of a direct beam stop just in front of our detector.
Several spent fuel rod samples were investigated.
Hot Cell Port
lead collimator
1.5 m
15 cm
7.5 cm
15 cm
Fuel Rod
Collimator
Lead
Collimator
Lead
Tungsten
Slit
First Mirror [ε,A]
Detector
Plane
Direct
beam stopper
8/30
Experimental Setup
ΦDet≃ ΦRod·ε·A
J. Ruz, February 6th 2024
Configuration 1
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 10
Detailed laboratory-based campaign conducted at the Irradiated Fuels Examination Laboratory (IFEL) at Oak Ridge National Laboratory.
The reflective optics were mounted downstream from the 30-inch-long lead collimator with an opening of 5mm x 69.85mm centered on the fuel rod. The 76-cm-long lead collimator was installed in the hot-cell port. An extra collimator with variable aperture was set between the optics and the lead collimation to minimize flux reaching the experimental setup. We also made use of a direct beam stop just in front of our detector.
Several spent fuel rod samples were investigated.
Hot Cell Port
lead collimator
9/30
Experimental Setup
1.5 m
15 cm
7.5 cm
15 cm
Fuel Rod
Collimator
Lead
Collimator
Lead
Tungsten
Slit
First Mirror [ε,A]
Detector
Plane
Direct
beam stopper
Second Mirror [ε,A]
ΦDet≃ ΦRod·ε2·A
J. Ruz, February 6th 2024
Configuration 2
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 11
10/30
Experimental Setup
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 12
To locate the reflected beam from the two-mirror reflective optic a GeGI detector was used. Unfortunately, its energy resolution did not allow for Pu to U ratio estimates due to a nearby Eu line.
A 5kg cryogenically cooled point contact coaxial High Purity Germanium (HPGe) was used for the purpose of the measurements.
To minimize background in the detector and improve the signal-to-noise ratio, the detectors were enclosed in ~ 5 cm of lead. Sn and Cu layers were included to reduce lead fluorescence.
Coaxial Germanium
GeGi Detector
11/30
Experimental setup
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 13
Experimental setup
Detectors
Efficiency simulations were carried out for the inverted coaxial detector using Geant4. We took measurements with a Ba-133 source to constrain the simulations. We characterize a dead layer of 0.565 ± 0.02 mm thick and efficiency curves for the coaxial germanium detector.
Coax dead layer (top) and efficiency curve (bottom)
Coaxial Germanium
Collimated 133Ba source
623 eV FWHM @ 81 keV
A cryogenically cooled point contact coaxial Hyper pure Germanium with rated FWHM of 700eV at 100keV was used for the purpose of the measurements.
12/30
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 14
MOX fuel rod
Detector
Experimental bay
In-wall Collimator
Measurement campaign
Ray-trace simulations
Measurements
  Direct beam
  Single reflection using M1(θ)
    MOX Center
    Improving s/n ratio by means of shielding
  Two reflections using M1 (θ) & M2 (3θ)
    MOX Center
    Ukα1 and Ukα2 rock
  Single reflection using M1(θ) & M2 (θ)
    MOX Center
13/30
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 15: 2D ray-tracing simulations enabled cross-checking of observables and guided adjustments in real-time
Ray-trace simulations
14/30
footprint
Single bounce
Emission spectra
Geometry
Reflectivity
Figure of error
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 16
Simulation M1@860arcsec & M2rock. Energy cut 91-100 keV
Kα1
Lower angle because higher energy
Kα2
Higher angle because lower energy
Double bounce
Ray-trace simulations
15/30
With M1 fix @ θ, tilting M2 @ 3 θ works as a spectrometer. We can select the energy of the throughput beam.
Just 0.016° !!
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 17: Characterize mirror performance
Measurement campaign
  Align system with position-sensitive, HPGe detector
  Generate “rocking curve”:
  flux vs mirror angle
    For Ka and Kb
16/30
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 18
GeGI measurements
Direct Beam Exposure
dwell 500s
Slit @ 500μm
17/30
Cumulative plot full energy range
Spectra
MOX is a very old fuel rod but still compromises the dead time of the detector
137Cs
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 19
GeGI measurements
Direct Beam Exposure
dwell 500s
Slit @ 500μm
18/30
Cumulative plot full energy range
Spectra
MOX is a very old fuel rod but still compromises the dead time of the detector
137Cs
Fluorescence Complex
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 20
GeGI measurements
M1 at 711 arcsec
dwell 1800s
Slit @ 700μm
19/30
Image full energy range
Cumulative events applying Energy cut for events between 91.5 and 103.0 keV
We can see at first glance a single reflection
Signal from reflection
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 21
GeGI measurements
M1 at 711 arcsec
dwell 1800s
Slit @ 700μm
20/30
Image full energy range
Cumulative events applying Energy cut for events between 91.5 and 103.0 keV
We can see at first glance a single reflection
Direct Beam
Reflected signal
Signal from reflection
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 22
GeGI measurements
M1 at 711 arcsec
21/30
Signal in the 90-102 keV energy range
Spectra comparison
θ=711 arcsec θlow background=711 arcsec
T=7200s Tlow background>24h
dwell xxx
Slit @ 1000μm
Pu and Eu
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 23
GeGI measurements
dwell >24hs
Slit @ 1000μm
22/30
M1 at 711 arcsec
Eu emission 105.30 keV
Pu-kα1 103.76 keV
∆E = 1.54 keV
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 24
Image U-kalpha1 Double Bounce
Image U-kalpha2 Double Bounce
M1 Single Bounce Leak
M1 Single Bounce Leak
M1&M2 Double Bounce
M1&M2 Double Bounce
GeGI measurements
dwell 1600s
Slit @ 1000μm
23/30
M1 at θ and M2 @ 3θ arcsec
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 25
M1 Single Bounce
U-kalpha1
Double Bounce
M1 Single Bounce
U-kalpha2
Double Bounce
Cumulative plot U-kalpha1
Cumulative plot U-kalpha2
GeGI measurements
dwell 1600s
Slit @ 1000μm
24/30
M1 at θ and M2 @ 3θ arcsec
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 26
GeGI measurements
25/30
M1 at θ and M2 @ 3θ arcsec
U-kalpha1 Single Bounce
U-kalpha2 Single Bounce
Rate 10%
of Single
Rate 9%
of Single
U-kalpha1
Double Bounce
Mean @2372 arcsec
U-kalpha2
Double Bounce
Mean @2446 arcsec
Δµ=74 arcsec
2 reflections
Rocking plot U-kalpha1 and Ukalpha2
1 reflection
Rocking U-kalpha1 and U-kalpha2
The two mirrors positioned at q /3q act as a fine spectrometer.
However, the rates for the stronger Ka lines are decreased by a factor of 10.
Equivalent to a prism mechanism to select different wavelengths
dwell 1600s
dwell 300s
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 27
Reflected image 90-102 keV energy range for the GeGI detector
GeGI measurements
M1 and M2 at 711 arcsec
dwell 1800
Slit @ 1000μm
Replace the GeGI detector by point contact coax
26/30
J. Ruz, February 6th 2024
The GeGI is an imaging detector useful to see the location of the beam
However, the energy resolution is insufficient to resolve the Pu Ka1 and Eu-155 peaks
It looks like the best strategy is to increase flux in a single reflection configuration while using a high energy resolution co-axial HPGe
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 28
Measurements COAX. MOX Center
dwell 28800s
Single Bounce with double surface
We observe a discrepancy of up to a 50%
dead time in the direct beam measurement as well as skewness of the fluorescence peaks
(not shown)
27/30
Spectra of reflected beam and background run
Eu emission 105.30 keV
Pu-kα1 103.76 keV
∆E = 1.54 keV
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 29
Bonus!! WAD-125 Center
28/30
Single Bounce with double surface
  Considerably younger fuel
We observe a discrepancy of up to a 60%
dead time in the direct beam measurement as well as skewness of the fluorescence peaks
(not shown)
Eu emission 105.30 keV
Pu-kα1 103.76 keV
∆E = 1.54 keV
Spectra of reflected beam and background run
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 30
29/30
Bonus!! WAD-125 Center vs. MOX
There is 134Cs in WAD-125
Half-life 134Cs is 2 years
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC

## Slide 31
Conclusions
30/30
The GeGI is an imaging detector useful to see the location of the beam but, the energy resolution is insufficient to resolve the Pu Ka1 and 155Eu peaks
Simulations and data agree that with two reflections we can to scan Ka1 and Ka2 but, the measured rate for two reflection is not sufficient to distinguish between Pu and 155Eu
Placing a coaxial point contact HPGe in the reflected beam (same location and shielding as for the GeGI) allows for much better energy resolution.
We are able to determine Pu/U ratio with very low statistical uncertainty. We could set the dwell based on the statistical uncertainty goal. For instance, with the MOX fuel a 0.5% statistical uncertainty will need of 600s dwell time.
J. Ruz, February 6th 2024
LLNL-PRES-859795
This work was performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under contract DE-AC52-07NA27344. Lawrence Livermore National Security, LLC