## Slide 1: Absolute calibration of monochromatic photons fluxes Application to the efficiency calibration of detectors
Marie-Christine Lépy, Yves Ménesguen, Víctor H. Elvíra
Laboratoire national Henri Becquerel (LNHB)
CEA Saclay - DRT/LIST/DIN/SIMRI/LNHB-MA

## Slide 2: Context - Metrology
To ensure quality, every measurement must be traceable to the International System of Units (SI) via reference standards.
The BIPM is the international organization established by the Metre Convention (1875), through which Member States act together on matters related to measurement science and measurement standards.
https://www.bipm.org/en/
Mission: international and impartial status to promote and advance the global comparability of measurements for:
scientific discovery and innovation,
industrial manufacturing and international trade,
improving the quality of life and sustaining the global environment.

## Slide 3: Context - Metrology
National Metrology Institutes (NMI) – intercomparisons - redundancy
LNE (Laboratoire National de métrologie et d’Essais): pilot of the French metrology
- several designated institutes for the different units
LNHB (Laboratoire national Henri Becquerel): in charge of ionising radiation
- radioactivity measurements (environment, nuclear power plants, radon, etc.)
- dosimetry (medical applications)
- caracterisation of radioactive and X-ray sources (X-ray tubes, electron beams, etc.)
- calibration of detectors (efficiency, response function)

## Slide 4: Outline
Cryogenic detector
BOLUX
Electrical substitution principle
Calibration setup
Measurement procedure
Correction factors
Results
Efficiency calibration of photodiodes
Efficiency calibration of energy-dispersive spectrometers
Application to the measurement of fluorescence yields
Conclusion & perspectives
Goal: determine the absolute photon flux intensity (s-1)
(independently from atomic or nuclear data)

## Slide 5: Cryogenic detectors
The incident x-ray beam in the absorber induces a temperature increase, ΔT, proportionnal to the incident energy
Continuous mode
Pulse mode
Application: spectroscopy
Application: bolometry
high-intensity flux measurement
T = 4K
T = ~ mK

## Slide 6: BOLUX - BOLometer for Use in the range of X-rays

## Slide 7: BOLUX - BOLometer for Use in the range of X-rays
Bolometer
Measurement chamber
Cryostat
Incoming flux
Collimators
Thermometer Pt1000
Thermometer cernox
bolometer
Wires thermalisation
Load resistor
Thermal screens

## Slide 8: Electrical substitution principle
Radiant power matched by electrical power
A thermistor with thermal link to the absorber is polarized (constant current)
The heating process (due to the radiant power) reduces the thermistor electrical resistance thus decreasing voltage.
The energy deposited by the radiation is known by finding the electrical power that needs to be transferred to the absorber in order to obtain the same temperature change (voltage).
η
Tension
Chauffage induit par rayonnement
Voltage
Voltage
Voltage
Beam open
Beam stopped

## Slide 9: Calibrating detectors
Photodiode
Measurement of the incident photons count rate
Semiconductor detector
Measurement of the detected count rate
Photodiode
Measurement of incident photons rate,
Semiconductor detector
Measurement of detected count rate,
Detection efficiency of the energy-dispersive spectrometer
+
BOLUX
Measurement of the beam power
Incident X-ray beam
Photodiode
Measurement of the current induced
Incident X-ray beam
1st step, with a high photon flux (~ 109 s-1)
The efficiency of both BOLUX and the photodiode depend on the photon energy
but neither of them are able to discriminate photons by energy
Monochromatic radiation
Efficiency (responsivity)
of the photodiode
Photodiode
Measurement of the current induced
BOLUX
Measurement of the beam power
+
2nd step, with a lower photon flux (~ 104 s-1)

## Slide 10: BOLUX : correction factor
Monte-Carlo simulations
Average energy deposited in the absorber per incident photon < 100 % due to:
Photons that do not interact with the absorber (transmission)
Photons that do interact but do not deposit all of their energy
  Characteristic X-ray escape
  Scattering
PENELOPE
PENetration and ENergy LOss of Positrons and Electrons *
* F. Salvat, OECD/NEA Data Bank (2015)
Énergie/keV
Energy (keV)

## Slide 11: Measurement results
Efficiency calibration of photodiodes
Energy (keV)
Calibration of the photodiodes vs BOLUX
Calibration of S8193 vs AXUV

## Slide 12: Measurement results
Efficiency calibration of energy-dispersive X-ray spectrometers - 1. SDD

## Slide 13: Measurement results
Efficiency calibration of energy-dispersive X-ray spectrometers - 2. HPGe
Energy (keV)
Efficiency
1st campaign
2nd campaign
Experimental results in the range 3.5 keV – 8 keV

## Slide 14: Measurement results
Monte Carlo simulations with different dead-layer thicknesses
Efficiency calibration of energy-dispersive X-ray spectrometers - 2. HPGe
Simulation – without dead layer

## Slide 15: Application
Reflection geometry: target at 45° both from the incident beam and the detector (SDD)
Incident monochromatic X-ray beam
Target
Photodiode
I0
It
Fluorescence
SDD
Sherman equation
Measurement of fluorescence yields

## Slide 16: Application
Example of measurement of K fluorescence yield: Fe

## Slide 17: Conclusion and perspectives
Absolute calibration of photon fluxes using an electrical-substitution cryogenic detector
Calibration (responsivity) of transfer photodiodes
Efficiency calibration of energy-dispersive spectrometers
Application to the measurement of fundamental parameters
  International Initiative on Fundamental Parameters
  https://www.exsa.hu/?inh=635
Application to the caracterisation of innovative materials
  XRR-GIXRF reference–free combined analysis
Extend the energy range of BOLUX with a thicker absorber (to reduce correction factors)

## Slide 18: Thank you for your attention
Acknowledgment:
SOLEIL and the staff of the METROLOGIE and PUMA beamlines
LNE for funding via project 258